## Overview
Sequence Overview
Sequence Details
Name: |  | 2025-02-24 EX3502 |  |  | Created On: | 2024-02-20 15:50
Directory: |  | Teaching methods\2025-EX3502 Lab |  |  | Created By: | s01yk2
Data Vault: |  | ChromeleonLocal |  |  | Updated On: | 2025-02-26 10:11
No. of Injections: |  | 49 |  |  | Updated By: | s01yk2
Injection Details
No. | Injection Name | Position | Type | Level | Inject Time | Status
1 | IPA S1 Repeatability check 1 | 1 | Unknown |  | 2025-02-25 15:20 | Finished
2 | IPA S1 Repeatability check 2 | 1 | Unknown |  | 2025-02-25 15:29 | Finished
3 | IPA S1 Repeatability check 3 | 1 | Unknown |  | 2025-02-25 15:38 | Finished
4 | Standard solution IPA S1 - 2.58M | 1 | Calibration Standard | 1 | 2025-02-25 15:47 | Finished
5 | Standard solution IPA S2 -3.93M | 2 | Calibration Standard | 2 | 2025-02-25 15:56 | Finished
6 | Standard solution IPA S3-5.16M | 3 | Calibration Standard | 3 | 2025-02-25 16:05 | Finished
7 | Standard solution IPA S4-6.53M | 4 | Calibration Standard | 4 | 2025-02-25 16:24 | Finished
8 | Standard solution IPA S5-7.71M | 5 | Calibration Standard | 5 | 2025-02-25 16:33 | Finished
9 | Standard solution IPA S6-9.11M | 6 | Calibration Standard | 6 | 2025-02-25 16:42 | Finished
10 | Standard solution IPA S7-10.36M | 7 | Calibration Standard | 7 | 2025-02-25 16:51 | Finished
11 | Standard solution IPA S8-11.67M | 8 | Calibration Standard | 8 | 2025-02-25 17:00 | Finished
12 | Standard solution IPA S9-12.44M | 1 | Calibration Standard | 9 | 2025-02-25 17:09 | Finished
13 | Standard solution IPA S10-12.9M | 2 | Calibration Standard | 10 | 2025-02-25 17:18 | Finished
14 | IPA S10 Repeatability check 1 | 2 | Unknown |  | 2025-02-25 17:27 | Finished
15 | IPA S10 Repeatability check 2 | 2 | Unknown |  | 2025-02-25 17:36 | Finished
16 | IPA S10 Repeatability check 3 | 2 | Unknown |  | 2025-02-25 17:46 | Finished
17 | G1_Sample 1 | 3 | Unknown |  | 2025-02-25 17:55 | Finished
18 | G1_Sample 1 | 3 | Unknown |  | 2025-02-25 18:04 | Finished
19 | G1_Sample 1 | 3 | Unknown |  | 2025-02-25 18:13 | Finished
20 | G1_Sample 2 | 4 | Unknown |  | 2025-02-25 18:22 | Finished
21 | G1_Sample 2 | 4 | Unknown |  | 2025-02-25 18:31 | Finished
22 | G1_Sample 2 | 4 | Unknown |  | 2025-02-25 18:40 | Finished
23 | G1_Sample 3 | 5 | Unknown |  | 2025-02-25 18:49 | Finished
24 | G1_Sample 3 | 5 | Unknown |  | 2025-02-25 18:58 | Finished
25 | G1_Sample 3 | 5 | Unknown |  | 2025-02-25 19:07 | Finished
26 | G1_Sample 4 | 6 | Unknown |  | 2025-02-25 19:16 | Finished
27 | G1_Sample 4 | 6 | Unknown |  | 2025-02-25 19:25 | Finished
28 | G1_Sample 4 | 6 | Unknown |  | 2025-02-25 19:35 | Finished
29 | G1_Sample 5 | 7 | Unknown |  | 2025-02-25 19:44 | Finished
30 | G1_Sample 5 | 7 | Unknown |  | 2025-02-25 19:53 | Finished
31 | G1_Sample 5 | 7 | Unknown |  | 2025-02-25 20:02 | Finished
32 | G1_Sample 6 | 8 | Unknown |  | 2025-02-25 20:11 | Finished
33 | G1_Sample 6 | 8 | Unknown |  | 2025-02-25 20:20 | Finished
34 | G1_Sample 6 | 8 | Unknown |  | 2025-02-25 20:29 | Finished
35 | G1_Sample 7 | 1 | Unknown |  | 2025-02-25 20:38 | Finished
36 | G1_Sample 7 | 1 | Unknown |  | 2025-02-25 20:47 | Finished
37 | G1_Sample 7 | 1 | Unknown |  | 2025-02-25 20:56 | Finished
38 | G1_Sample 8 | 2 | Unknown |  | 2025-02-25 21:05 | Finished
39 | G1_Sample 8 | 2 | Unknown |  | 2025-02-25 21:15 | Finished
40 | G1_Sample 8 | 2 | Unknown |  | 2025-02-25 21:24 | Finished
41 | G1_Sample 9 | 3 | Unknown |  | 2025-02-25 21:33 | Finished
42 | G1_Sample 9 | 3 | Unknown |  | 2025-02-25 21:42 | Finished
43 | G1_Sample 9 | 3 | Unknown |  | 2025-02-25 21:51 | Finished
44 | G1_Sample 10 | 4 | Unknown |  | 2025-02-25 22:00 | Finished
45 | G1_Sample 10 | 4 | Unknown |  | 2025-02-25 22:09 | Finished
46 | G1_Sample 10 | 4 | Unknown |  | 2025-02-25 22:18 | Finished
47 | IPA 10% initial solution for boiler | 5 | Unknown |  | 2025-02-25 22:27 | Finished
48 | IPA 10% initial solution for boiler | 5 | Unknown |  | 2025-02-25 22:36 | Finished
49 | IPA 10% initial solution for boiler | 5 | Unknown |  | 2025-02-25 22:45 | Finished

## Summary
Summary
Sequence Details
Name: |  | 2025-02-24 EX3502 |  |  | Created On: | 2024-02-20 15:50
Directory: |  | Teaching methods\2025-EX3502 Lab |  |  | Created By: | s01yk2
Data Vault: |  | ChromeleonLocal |  |  | Updated On: | 2025-02-26 10:11
No. of Injections: |  | 49 |  |  | Updated By: | s01yk2
By Component |  | IPA
No. | Injection Name | Ret.Time | Area | Height | Amount | Rel.Area | Peak Type | %Mol of IPA
|  | min | pA*min | pA | mol/l | %
|  | FID | FID | FID | FID | FID | FID
|  | IPA | IPA | IPA | IPA | IPA | IPA
1 | IPA S1 Repeatability check 1 | 1.7933 | 8.9257 | 201.842 | 2.6948 | 100 | BMB | 5.768%
2 | IPA S1 Repeatability check 2 | 1.7933 | 8.7749 | 205.648 | 2.6585 | 100 | BMB | 5.676%
3 | IPA S1 Repeatability check 3 | 1.795 | 8.8086 | 200.1461 | 2.6667 | 100 | BMB | 5.696%
4 | Standard solution IPA S1 - 2.58M | 1.7967 | 8.8055 | 202.6729 | 2.6659 | 100 | BMB* | 5.694%
5 | Standard solution IPA S2 -3.93M | 1.795 | 14.0529 | 336.107 | 3.9283 | 100 | BMB* | 9.194%
6 | Standard solution IPA S3-5.16M | 1.8 | 19.223 | 449.5344 | 5.1721 | 100 | BMB | 13.37%
7 | Standard solution IPA S4-6.53M | 1.8 | 24.5232 | 572.1431 | 6.4472 | 100 | BMB | 18.65%
8 | Standard solution IPA S5-7.71M | 1.805 | 29.8013 | 689.3574 | 7.7169 | 100 | BMB | 25.34%
9 | Standard solution IPA S6-9.11M | 1.805 | 35.1112 | 815.1168 | 8.9944 | 100 | BMB | 34.18%
10 | Standard solution IPA S7-10.36M | 1.8083 | 40.9471 | 937.8092 | 10.3983 | 100 | BMB* | 47.78%
11 | Standard solution IPA S8-11.67M | 1.81 | 45.6857 | 1053.724 | 11.5383 | 100 | BMB* | 63.86%
12 | Standard solution IPA S9-12.44M | 1.8133 | 49.7034 | 1128.6904 | 12.5049 | 100 | BMB* | 83.72%
13 | Standard solution IPA S10-12.9M | 1.815 | 51.86 | 1165.9821 | 13.0237 | 100 | BM * | 98.26%
14 | IPA S10 Repeatability check 1 | 1.8133 | 51.5503 | 1155.2229 | 12.9492 | 100 | BMB* | 95.95%
15 | IPA S10 Repeatability check 2 | 1.8133 | 51.879 | 1171.5981 | 13.0283 | 100 | BMB* | 98.4%
16 | IPA S10 Repeatability check 3 | 1.8133 | 50.8009 | 1161.684 | 12.7689 | 100 | BMB* | 90.69%
17 | G1_Sample 1 | 1.7967 | 14.5146 | 322.7372 | 4.0394 | 100 | BMB | 9.534%
18 | G1_Sample 1 | 1.7967 | 14.3524 | 321.8536 | 4.0003 | 100 | BMB | 9.414%
19 | G1_Sample 1 | 1.7967 | 14.4909 | 321.697 | 4.0337 | 100 | BMB* | 9.517%
20 | G1_Sample 2 | 1.795 | 13.7445 | 313.6846 | 3.8541 | 100 | BMB* | 8.97%
21 | G1_Sample 2 | 1.795 | 13.7872 | 301.6403 | 3.8644 | 100 | BMB* | 9.001%
22 | G1_Sample 2 | 1.7933 | 13.7136 | 312.7292 | 3.8467 | 100 | BMB* | 8.948%
23 | G1_Sample 3 | 1.81 | 44.8085 | 1006.1113 | 11.3273 | 100 | BMB* | 60.41%
24 | G1_Sample 3 | 1.81 | 44.0229 | 1005.5526 | 11.1383 | 100 | BMB* | 57.52%
25 | G1_Sample 3 | 1.81 | 44.61 | 1006.2692 | 11.2795 | 100 | BMB* | 59.66%
26 | G1_Sample 4 | 1.7967 | 13.0273 | 277.302 | 3.6816 | 100 | BMB* | 8.458%
27 | G1_Sample 4 | 1.795 | 13.018 | 274.981 | 3.6793 | 100 | BMB* | 8.452%
28 | G1_Sample 4 | 1.7967 | 13.0459 | 282.8655 | 3.686 | 100 | BMB* | 8.471%
29 | G1_Sample 5 | 1.81 | 44.758 | 991.6679 | 11.3151 | 100 | BMB* | 60.22%
30 | G1_Sample 5 | 1.81 | 44.7794 | 1012.9053 | 11.3203 | 100 | BMB* | 60.3%
31 | G1_Sample 5 | 1.8083 | 44.2703 | 981.2505 | 11.1978 | 100 | BMB* | 58.41%
32 | G1_Sample 6 | 1.795 | 12.6874 | 277.3647 | 3.5998 | 100 | BMB* | 8.22%
33 | G1_Sample 6 | 1.7967 | 12.6889 | 274.1567 | 3.6002 | 100 | BMB | 8.221%
34 | G1_Sample 6 | 1.795 | 12.5842 | 276.0927 | 3.575 | 100 | BMB | 8.149%
35 | G1_Sample 7 | 1.7967 | 12.565 | 261.9027 | 3.5703 | 100 | BMB | 8.135%
36 | G1_Sample 7 | 1.795 | 12.4719 | 268.8878 | 3.5479 | 100 | BMB* | 8.071%
37 | G1_Sample 7 | 1.795 | 12.496 | 262.5258 | 3.5537 | 100 | BMB | 8.088%
38 | G1_Sample 8 | 1.8067 | 44.6952 | 982.7082 | 11.3 | 100 | BMB* | 59.98%
39 | G1_Sample 8 | 1.8083 | 43.9645 | 968.1181 | 11.1242 | 100 | BMB* | 57.31%
40 | G1_Sample 8 | 1.8067 | 44.5477 | 985.0626 | 11.2645 | 100 | BMB* | 59.43%
41 | G1_Sample 9 | 1.795 | 12.4841 | 273.8579 | 3.5509 | 100 | BMB | 8.079%
42 | G1_Sample 9 | 1.795 | 12.4565 | 266.4261 | 3.5442 | 100 | BMB | 8.06%
43 | G1_Sample 9 | 1.795 | 12.4598 | 268.4485 | 3.545 | 100 | BMB | 8.063%
44 | G1_Sample 10 | 1.8083 | 44.6076 | 971.2986 | 11.279 | 100 | BMB* | 59.65%
45 | G1_Sample 10 | 1.8083 | 44.6701 | 984.3211 | 11.294 | 100 | BMB* | 59.88%
46 | G1_Sample 10 | 1.8067 | 44.0119 | 966.2205 | 11.1356 | 100 | BMB | 57.48%
47 | IPA 10% initial solution for boiler | 1.795 | 14.7617 | 325.4839 | 4.0988 | 100 | BMB | 9.719%
48 | IPA 10% initial solution for boiler | 1.795 | 14.7773 | 317.434 | 4.1026 | 100 | BMB | 9.73%
49 | IPA 10% initial solution for boiler | 1.7933 | 14.5547 | 326.6239 | 4.049 | 100 | BMB* | 9.564%

## Calibration
Calibration
Calibration Details |  | IPA
Calibration Type |  | Lin, WithOffset, Avg |  |  | Offset (C0) |  | -2.2759
Evaluation Type |  | Area |  |  | Slope (C1) |  | 4.1567
Number of Calibration Points |  | 10 |  |  | Curve (C2) |  | 0
Number of disabled Calibration Points |  | 0 |  |  | R-Square |  | 0.99945
Calibration Plot |  | IPA
Calibration Results |  | IPA
No. | Injection Name | Calibration | X Value | Y Value | Y Value | Area | Height
|  | Level |  |  |  | pA*min | pA
|  |  | FID | FID | FID | FID | FID
|  |  | IPA | IPA | IPA | IPA | IPA
4 | Standard solution IPA S1 - 2.58M | 1 | 2.58 | 8.8055 | 8.8055 | 8.8055 | 202.6729
5 | Standard solution IPA S2 -3.93M | 2 | 3.93 | 14.0529 | 14.0529 | 14.0529 | 336.107
6 | Standard solution IPA S3-5.16M | 3 | 5.16 | 19.223 | 19.223 | 19.223 | 449.5344
7 | Standard solution IPA S4-6.53M | 4 | 6.53 | 24.5232 | 24.5232 | 24.5232 | 572.1431
8 | Standard solution IPA S5-7.71M | 5 | 7.71 | 29.8013 | 29.8013 | 29.8013 | 689.3574
9 | Standard solution IPA S6-9.11M | 6 | 9.11 | 35.1112 | 35.1112 | 35.1112 | 815.1168
10 | Standard solution IPA S7-10.36M | 7 | 10.36 | 40.9471 | 40.9471 | 40.9471 | 937.8092
11 | Standard solution IPA S8-11.67M | 8 | 11.67 | 45.6857 | 45.6857 | 45.6857 | 1053.724
12 | Standard solution IPA S9-12.44M | 9 | 12.44 | 49.7034 | 49.7034 | 49.7034 | 1128.6904
13 | Standard solution IPA S10-12.9M | 10 | 12.9 | 51.86 | 51.86 | 51.86 | 1165.9821

## Chromatogram
Chromatogram

## Audit Trail
Audit Trail
Injection Details
Injection Name: |  | IPA S1 Repeatability check 1 | Run Time (min): | 3.6983
Vial Number: |  | 1 | Injection Volume: | 0.5
Injection Type: |  | Unknown | Channel: | FID
Calibration Level: |  |  | Wavelength: | n.a.
Instrument Method: |  | IPA Stabilwax 2024-02-20_high conc | Bandwidth: | n.a.
Processing Method: |  | IPA_2024-02-19_high conc | Dilution Factor: | 1
Injection Date/Time: |  | 2025-02-25 15:20 | Sample Weight: | 1
Day Time | Ret. Time | Command/Message
(UTC+00:00) | min
15:15:12 |  | Audit trail of injection "chrom://md-051298/ChromeleonLocal/Teaching methods/2025-EX3502 Lab/2025-02-24 EX3502.seq/829.smp". Executed 25/02/2025 on instrument TRACE1300 (instrument controller MD-051298).
15:15:12 |  | Start of injection 12 "IPA S1", using instrument method "IPA Stabilwax 2024-02-20_high conc".
15:15:12 | 0.000 | Entered stage "Instrument Setup"
15:15:12 | 0.000 | GC.SSL.FlowMode = FlowCtrl
15:15:12 | 0.000 | Sampler.DrawSpeed = Slow
15:15:12 | 0.000 | Sampler.FillStrokes = 3
15:15:12 | 0.000 | Sampler.AirVolume = 1.00 µl
15:15:12 | 0.000 | Sampler.GcType = TRACE_1300_1310
15:15:12 | 0.000 | Sampler.PostWash = 6
15:15:12 | 0.000 | Sampler.PostWashVial = D
15:15:12 | 0.000 | Sampler.SampleWash = 3
15:15:12 | 0.000 | Sampler.PreWash = 5
15:15:12 | 0.000 | Sampler.PreWashVial = B
15:15:12 | 0.000 | GC.PrepRunTimeout = 999.99 min
15:15:12 | 0.000 | GC.EquilibrationTime = 0.50 min
15:15:12 | 0.000 | GC.ReadyDelay = 0.00 min
15:15:12 | 0.000 | GC.SSL.SplitMode = Split
15:15:12 | 0.000 | GC.SSL.SplitFlow.Nominal = 240.0 ml/min
15:15:12 | 0.000 | GC.SSL.SplitFlowCtrl = On
15:15:12 | 0.000 | GC.SSL.PurgeFlow.Nominal = 5.000 ml/min
15:15:12 | 0.000 | GC.SSL.PurgeFlowCtrl = On
15:15:12 | 0.000 | GC.SSL.VacuumCorrection = Off
15:15:12 | 0.000 | GC.SSL.GasSaverFlow = 10.0 ml/min
15:15:12 | 0.000 | GC.SSL.GasSaverTime = 2.00 min
15:15:12 | 0.000 | GC.SSL.GasSaverCtrl = On
15:15:12 | 0.000 | GC.SSL.Temperature.Nominal = 180 °C
15:15:12 | 0.000 | GC.SSL.TempCtrl = On
15:15:12 | 0.000 | GC.FrontColumn.Description = "Stabilwax-DA SN1645029"
15:15:12 | 0.000 | GC.FrontColumn.Length = 30.00 m
15:15:12 | 0.000 | GC.FrontColumn.NominalID = 0.250 mm
15:15:12 | 0.000 | GC.FrontColumn.FilmThickness = 0.25 µm
15:15:12 | 0.000 | GC.FID.Data_Collection_Rate = 10 Hz
15:15:12 | 0.000 | GC.FID.AirFlow.Nominal = 350.0 ml/min
15:15:12 | 0.000 | GC.FID.AirFlowCtrl = On
15:15:12 | 0.000 | GC.FID.MakeupFlowCtrl = Off
15:15:12 | 0.000 | GC.FID.H2Flow.Nominal = 35.0 ml/min
15:15:12 | 0.000 | GC.FID.H2FlowCtrl = On
15:15:12 | 0.000 | GC.FID.Temperature.Nominal = 260 °C
15:15:12 | 0.000 | GC.FID.TempCtrl = On
15:15:12 | 0.000 | GC.FID.Flame = On
15:15:12 | 0.000 | GC.FID.FlameoutRetry = On
15:15:12 | 0.000 | GC.FID.LitOffset = 1.0 pA
15:15:12 | 0.000 | GC.FID.PeakWidth = Standard
15:15:12 | 0.000 | GC.SSL.FlowCtrl = On
15:15:12 | 0.000 | GC.TempCtrl = On
15:15:12 | 0.000 | Entered stage "Equilibration"
15:15:12 | 0.000 | GC.SSL.Flow.Nominal = 1.200 ml/min
15:15:12 | 0.000 | GC.Temperature.Nominal = 40.0 °C
15:15:12 | 0.000 | Entered stage "Inject Preparation"
15:15:12 | 0.000 | GC.FID.Autozero
15:15:12 | 0.000 | Wait Sampler.Ready And GC.Ready
15:16:14 | 0.000 | New detailed run state: Waiting for equilibration time.
15:16:44 | 0.000 | Log RunState: StandBy
15:16:44 | 0.000 | Prep-run command sent to GC.
15:16:44 | 0.000 | Log RunState: PrepRun
15:16:44 | 0.000 | New detailed run state: Executing prep-run events
15:17:05 | 0.000 | New detailed run state: Waiting for start key
15:17:05 | 0.000 | Log RunState: Ready
15:17:05 | 0.000 | Log Ready: Ready
15:17:05 | 0.000 | Wait finished
15:17:05 | 0.000 | Entered stage "Inject"
15:17:05 | 0.000 | Sampler.Inject
15:17:05 | 0.000 | Injecting from vial position 1.
15:17:05 | 0.000 | Injection Volume is 0.5 µl.
15:17:05 | 0.000 | Waiting for inject response on Sampler.
15:19:52 | 0.000 | Got inject response.
15:19:52 | 0.000 | Entered stage "Start Run"
15:19:52 | 0.000 | GC.FID.AcqOn
15:19:52 | 0.000 | Entered stage "Run"
15:19:53 | 0.005 | New detailed run state: Running
15:19:53 | 0.005 | Log RunState: Run
15:19:53 | 0.005 | Log Ready: NotReady
15:20:52 | 1.000 | GC.Temperature.Nominal Gradient Start = 40.0 °C, End = 125.0 °C, Duration = 1.700 min
15:22:34 | 2.700 | GC.Temperature.Nominal = 125.0 °C
15:23:34 | 3.700 | GC.Temperature.Nominal = 125.0 °C
15:23:34 | 3.700 | Entered stage "Stop Run"
15:23:34 | 3.700 | GC.FID.AcqOff
15:23:34 | 3.700 | GC.SSL.Flow.Nominal = 1.200 ml/min
15:23:34 | 3.700 | Terminating instrument method chrom://md-051298/ChromeleonLocal/Teaching methods/2025-EX3502 Lab/2025-02-24 EX3502.seq/IPA Stabilwax 2024-02-20_high conc.instmeth. Waiting for instrument to finish processing of current injection ...
15:23:40 | 3.791 | Log RunState: Cool
15:23:40 | 3.793 | New detailed run state: Cooling down
15:23:46 | 3.889 | Log RunState: NotReady
15:23:46 | 3.889 | OK to terminate instrument method chrom://md-051298/ChromeleonLocal/Teaching methods/2025-EX3502 Lab/2025-02-24 EX3502.seq/IPA Stabilwax 2024-02-20_high conc.instmeth.
15:23:46 |  | End of injection "IPA S1".

## Standards
2025-02-24
| Total Volume = 5 ml |  | Exact masses
Standard solution name | V(IPA), ml | V(H2O), ml | m (IPA), g | m (H2O), g | n (IPA), mol | n (H2O), mol | % Mol | C(IPA), Mol/L
Standard solution IPA S1 | 1 | 4 | 0.7757 | 4.0047 | 0.012907 | 0.22248 | 5.483% | 2.5814
Standard solution IPA S2 | 1.5 | 3.5 | 1.1814 | 3.5057 | 0.019657 | 0.19476 | 9.168% | 3.9314
Standard solution IPA S3 | 2 | 3 | 1.5518 | 3.0111 | 0.02582 | 0.16728 | 13.37% | 5.1641
Standard solution IPA S4 | 2.5 | 2.5 | 1.9612 | 2.5184 | 0.032632 | 0.13991 | 18.91% | 6.5265
Standard solution IPA S5 | 3 | 2 | 2.3177 | 2.0137 | 0.038564 | 0.11187 | 25.63% | 7.7128
Standard solution IPA S6 | 3.5 | 1.5 | 2.7382 | 1.5146 | 0.045561 | 0.084144 | 35.13% | 9.1121
Standard solution IPA S7 | 4 | 1 | 3.1122 | 0.9931 | 0.051784 | 0.055172 | 48.42% | 10.3567
Standard solution IPA S8 | 4.5 | 0.5 | 3.5072 | 0.4961 | 0.058356 | 0.027561 | 67.92% | 11.6712
Standard solution IPA S9 | 1.8 | 0.2 | 3.7392 | 0.1975 | 0.062216 | 0.010972 | 85.01% | 12.4433
Standard solution IPA S10 | 5 | 0 | 3.8769 | 0 | 0.064507 | 0 | 100% | 12.9015
MW IPA, g/mol | MW H2O, g/mol | IPA Density @ 20°C, g/mL | Water Density @ 20°C, g/mL
60.1 | 18 | 0.786 | 0.9982