## Overview
Sequence Overview
Sequence Details
Name: |  | 2025-02-27 10 am EX3502 |  |  | Created On: | 2024-02-20 15:50
Directory: |  | Teaching methods\2025-EX3502 Lab |  |  | Created By: | s01yk2
Data Vault: |  | ChromeleonLocal |  |  | Updated On: | 2025-02-28 09:06
No. of Injections: |  | 90 |  |  | Updated By: | s01yk2
Injection Details
No. | Injection Name | Position | Type | Level | Inject Time | Status
1 | IPA S1 Repeatability check 1 | 1 | Unknown |  | 2025-02-27 09:44 | Finished
2 | IPA S1 Repeatability check 2 | 1 | Unknown |  | 2025-02-27 09:53 | Finished
3 | IPA S1 Repeatability check 3 | 1 | Unknown |  | 2025-02-27 10:02 | Finished
4 | Standard solution IPA S1 - 2.58M | 1 | Calibration Standard | 1 | 2025-02-27 10:11 | Finished
5 | Standard solution IPA S2 -3.93M | 2 | Calibration Standard | 2 | 2025-02-27 10:20 | Finished
6 | Standard solution IPA S3-5.16M | 3 | Calibration Standard | 3 | 2025-02-27 10:29 | Finished
7 | Standard solution IPA S4-6.53M | 4 | Calibration Standard | 4 | 2025-02-27 10:38 | Finished
8 | Standard solution IPA S5-7.71M | 5 | Calibration Standard | 5 | 2025-02-27 10:47 | Finished
9 | Standard solution IPA S6-9.11M | 6 | Calibration Standard | 6 | 2025-02-27 10:56 | Finished
10 | Standard solution IPA S7-10.36M | 7 | Calibration Standard | 7 | 2025-02-27 11:04 | Finished
11 | Standard solution IPA S8-11.67M | 8 | Calibration Standard | 8 | 2025-02-27 11:13 | Finished
12 | Standard solution IPA S9-12.44M | 1 | Calibration Standard | 9 | 2025-02-27 11:22 | Finished
13 | Standard solution IPA S10-12.9M | 2 | Calibration Standard | 10 | 2025-02-27 11:31 | Finished
14 | IPA S10 Repeatability check 1 | 2 | Unknown |  | 2025-02-27 11:40 | Finished
15 | IPA S10 Repeatability check 2 | 2 | Unknown |  | 2025-02-27 11:49 | Finished
16 | IPA S10 Repeatability check 3 | 2 | Unknown |  | 2025-02-27 11:58 | Finished
17 | IPA 10% initial solution for boiler | 3 | Unknown |  | 2025-02-27 12:07 | Finished
18 | IPA 10% initial solution for boiler | 3 | Unknown |  | 2025-02-27 12:16 | Finished
19 | G2_Sample 1 | 1 | Unknown |  | 2025-02-27 12:25 | Finished
20 | G2_Sample 1 | 1 | Unknown |  | 2025-02-27 12:33 | Finished
21 | G2_Sample 1 | 1 | Unknown |  | 2025-02-27 12:42 | Finished
22 | G2_Sample 2 | 2 | Unknown |  | 2025-02-27 12:51 | Finished
23 | G2_Sample 2 | 2 | Unknown |  | 2025-02-27 13:00 | Finished
24 | G2_Sample 2 | 2 | Unknown |  | 2025-02-27 13:09 | Finished
25 | G2_Sample 3 | 3 | Unknown |  | 2025-02-27 13:18 | Finished
26 | G2_Sample 3 | 3 | Unknown |  | 2025-02-27 13:27 | Finished
27 | G2_Sample 3 | 3 | Unknown |  | 2025-02-27 13:36 | Finished
28 | G2_Sample 4 | 4 | Unknown |  | 2025-02-27 13:45 | Finished
29 | G2_Sample 4 | 4 | Unknown |  | 2025-02-27 13:54 | Finished
30 | G2_Sample 4 | 4 | Unknown |  | 2025-02-27 14:03 | Finished
31 | G2_Sample 5 | 5 | Unknown |  | 2025-02-27 14:12 | Finished
32 | G2_Sample 5 | 5 | Unknown |  | 2025-02-27 14:21 | Finished
33 | G2_Sample 5 | 5 | Unknown |  | 2025-02-27 14:30 | Finished
34 | G2_Sample 6 | 6 | Unknown |  | 2025-02-27 14:39 | Finished
35 | G2_Sample 6 | 6 | Unknown |  | 2025-02-27 14:48 | Finished
36 | G2_Sample 6 | 6 | Unknown |  | 2025-02-27 14:57 | Finished
37 | G2_Sample 7 | 7 | Unknown |  | 2025-02-27 15:06 | Finished
38 | G2_Sample 7 | 7 | Unknown |  | 2025-02-27 15:15 | Finished
39 | G2_Sample 7 | 7 | Unknown |  | 2025-02-27 15:24 | Finished
40 | G2_Sample 8 | 8 | Unknown |  | 2025-02-27 15:33 | Finished
41 | G2_Sample 8 | 8 | Unknown |  | 2025-02-27 15:42 | Finished
42 | G2_Sample 8 | 8 | Unknown |  | 2025-02-27 15:51 | Finished
43 | G2_Sample 9 | 1 | Unknown |  | 2025-02-27 16:00 | Finished
44 | G2_Sample 9 | 1 | Unknown |  | 2025-02-27 16:09 | Finished
45 | G2_Sample 9 | 1 | Unknown |  | 2025-02-27 16:18 | Finished
46 | G2_Sample 10 | 2 | Unknown |  | 2025-02-27 16:27 | Finished
47 | G2_Sample 10 | 2 | Unknown |  | 2025-02-27 16:36 | Finished
48 | G2_Sample 10 | 2 | Unknown |  | 2025-02-27 16:45 | Finished
49 | G2_Sample 11 | 3 | Unknown |  | 2025-02-27 16:54 | Finished
50 | G2_Sample 11 | 3 | Unknown |  | 2025-02-27 17:03 | Finished
51 | G2_Sample 11 | 3 | Unknown |  | 2025-02-27 17:12 | Finished
52 | G2_Sample 12 | 4 | Unknown |  | 2025-02-27 17:21 | Finished
53 | G2_Sample 12 | 4 | Unknown |  | 2025-02-27 17:30 | Finished
54 | G2_Sample 12 | 4 | Unknown |  | 2025-02-27 17:39 | Finished
55 | G3_Sample 1 | 5 | Unknown |  | 2025-02-27 17:48 | Finished
56 | G3_Sample 1 | 5 | Unknown |  | 2025-02-27 17:57 | Finished
57 | G3_Sample 1 | 5 | Unknown |  | 2025-02-27 18:06 | Finished
58 | G3_Sample 2 | 6 | Unknown |  | 2025-02-27 18:15 | Finished
59 | G3_Sample 2 | 6 | Unknown |  | 2025-02-27 18:24 | Finished
60 | G3_Sample 2 | 6 | Unknown |  | 2025-02-27 18:33 | Finished
61 | G3_Sample 3 | 7 | Unknown |  | 2025-02-27 18:42 | Finished
62 | G3_Sample 3 | 7 | Unknown |  | 2025-02-27 18:51 | Finished
63 | G3_Sample 3 | 7 | Unknown |  | 2025-02-27 19:00 | Finished
64 | G3_Sample 4 | 8 | Unknown |  | 2025-02-27 19:10 | Finished
65 | G3_Sample 4 | 8 | Unknown |  | 2025-02-27 19:19 | Finished
66 | G3_Sample 4 | 8 | Unknown |  | 2025-02-27 19:28 | Finished
67 | G3_Sample 5 | 1 | Unknown |  | 2025-02-27 19:37 | Finished
68 | G3_Sample 5 | 1 | Unknown |  | 2025-02-27 19:46 | Finished
69 | G3_Sample 5 | 1 | Unknown |  | 2025-02-27 19:55 | Finished
70 | G3_Sample 6 | 2 | Unknown |  | 2025-02-27 20:04 | Finished
71 | G3_Sample 6 | 2 | Unknown |  | 2025-02-27 20:13 | Finished
72 | G3_Sample 6 | 2 | Unknown |  | 2025-02-27 20:22 | Finished
73 | G3_Sample 7 | 3 | Unknown |  | 2025-02-27 20:30 | Finished
74 | G3_Sample 7 | 3 | Unknown |  | 2025-02-27 20:39 | Finished
75 | G3_Sample 7 | 3 | Unknown |  | 2025-02-27 20:48 | Finished
76 | G3_Sample 8 | 4 | Unknown |  | 2025-02-27 20:57 | Finished
77 | G3_Sample 8 | 4 | Unknown |  | 2025-02-27 21:06 | Finished
78 | G3_Sample 8 | 4 | Unknown |  | 2025-02-27 21:15 | Finished
79 | G3_Sample 9 | 5 | Unknown |  | 2025-02-27 21:24 | Finished
80 | G3_Sample 9 | 5 | Unknown |  | 2025-02-27 21:33 | Finished
81 | G3_Sample 9 | 5 | Unknown |  | 2025-02-27 21:42 | Finished
82 | G3_Sample 10 | 6 | Unknown |  | 2025-02-27 21:51 | Finished
83 | G3_Sample 10 | 6 | Unknown |  | 2025-02-27 22:00 | Finished
84 | G3_Sample 10 | 6 | Unknown |  | 2025-02-27 22:09 | Finished
85 | G3_Sample 11 | 7 | Unknown |  | 2025-02-27 22:18 | Finished
86 | G3_Sample 11 | 7 | Unknown |  | 2025-02-27 22:27 | Finished
87 | G3_Sample 11 | 7 | Unknown |  | 2025-02-27 22:36 | Finished
88 | G3_Sample 12 | 8 | Unknown |  | 2025-02-27 22:45 | Finished
89 | G3_Sample 12 | 8 | Unknown |  | 2025-02-27 22:54 | Finished
90 | G3_Sample 12 | 8 | Unknown |  | 2025-02-27 23:03 | Finished

## Summary
Summary
Sequence Details
Name: |  | 2025-02-27 10 am EX3502 |  |  | Created On: | 2024-02-20 15:50
Directory: |  | Teaching methods\2025-EX3502 Lab |  |  | Created By: | s01yk2
Data Vault: |  | ChromeleonLocal |  |  | Updated On: | 2025-02-28 09:06
No. of Injections: |  | 90 |  |  | Updated By: | s01yk2
By Component |  | IPA
No. | Injection Name | Ret.Time | Area | Height | Amount | Rel.Area | Peak Type | %Mol of IPA
|  | min | pA*min | pA | mol/l | %
|  | FID | FID | FID | FID | FID | FID
|  | IPA | IPA | IPA | IPA | IPA | IPA
1 | IPA S1 Repeatability check 1 | 1.7933 | 8.3546 | 202.0271 | 2.6045 | 100 | BMB | 5.54%
2 | IPA S1 Repeatability check 2 | 1.7917 | 8.6703 | 212.7999 | 2.6818 | 100 | BMB | 5.735%
3 | IPA S1 Repeatability check 3 | 1.7917 | 8.6457 | 201.3159 | 2.6758 | 100 | BMB | 5.719%
4 | Standard solution IPA S1 - 2.58M | 1.7917 | 8.6425 | 206.4439 | 2.675 | 100 | BMB | 5.717%
5 | Standard solution IPA S2 -3.93M | 1.7933 | 13.7896 | 327.2576 | 3.9348 | 100 | BMB | 9.214%
6 | Standard solution IPA S3-5.16M | 1.795 | 18.8365 | 447.4335 | 5.17 | 100 | BMB | 13.36%
7 | Standard solution IPA S4-6.53M | 1.7983 | 24.1656 | 567.6417 | 6.4743 | 100 | BMB | 18.78%
8 | Standard solution IPA S5-7.71M | 1.7983 | 29.1615 | 676.3403 | 7.6971 | 100 | BMB | 25.22%
9 | Standard solution IPA S6-9.11M | 1.8017 | 34.4926 | 793.8218 | 9.0019 | 100 | BMB | 34.25%
10 | Standard solution IPA S7-10.36M | 1.805 | 39.5562 | 912.1902 | 10.2413 | 100 | BMB* | 45.99%
11 | Standard solution IPA S8-11.67M | 1.805 | 45.296 | 1020.839 | 11.6461 | 100 | BMB | 65.73%
12 | Standard solution IPA S9-12.44M | 1.8083 | 49.0797 | 1114.0516 | 12.5722 | 100 | BMB* | 85.42%
13 | Standard solution IPA S10-12.9M | 1.8083 | 50.7344 | 1132.6117 | 12.9772 | 100 | BMB* | 96.8%
14 | IPA S10 Repeatability check 1 | 1.8083 | 50.829 | 1146.0485 | 13.0003 | 100 | BMB* | 97.52%
15 | IPA S10 Repeatability check 2 | 1.8067 | 50.6684 | 1120.0943 | 12.961 | 100 | BMB* | 96.31%
16 | IPA S10 Repeatability check 3 | 1.8083 | 50.6406 | 1099.8564 | 12.9542 | 100 | BMB* | 96.1%
17 | IPA 10% initial solution for boiler | 1.7933 | 14.3868 | 318.3137 | 4.081 | 100 | BMB | 9.663%
18 | IPA 10% initial solution for boiler | 1.7917 | 14.3906 | 319.3706 | 4.0819 | 100 | BMB | 9.666%
19 | G2_Sample 1 | 1.7917 | 12.1965 | 266.8956 | 3.5449 | 100 | BMB | 8.062%
20 | G2_Sample 1 | 1.7917 | 12.1515 | 268.3636 | 3.5339 | 100 | BMB | 8.031%
21 | G2_Sample 1 | 1.79 | 12.1928 | 264.2236 | 3.544 | 100 | BMB | 8.06%
22 | G2_Sample 2 | 1.805 | 42.3414 | 941.7973 | 10.923 | 100 | BMB* | 54.45%
23 | G2_Sample 2 | 1.805 | 43.4655 | 957.3504 | 11.1981 | 100 | BMB* | 58.41%
24 | G2_Sample 2 | 1.805 | 43.4696 | 967.6879 | 11.1991 | 100 | BMB* | 58.43%
25 | G2_Sample 3 | 1.79 | 12.3312 | 263.5116 | 3.5778 | 100 | BMB | 8.157%
26 | G2_Sample 3 | 1.79 | 12.0943 | 261.6133 | 3.5198 | 100 | BMB | 7.991%
27 | G2_Sample 3 | 1.7917 | 12.1428 | 256.063 | 3.5317 | 100 | BMB | 8.024%
28 | G2_Sample 4 | 1.805 | 43.0305 | 951.2216 | 11.0916 | 100 | BMB* | 56.84%
29 | G2_Sample 4 | 1.805 | 43.551 | 940.8007 | 11.219 | 100 | BMB* | 58.73%
30 | G2_Sample 4 | 1.805 | 42.8138 | 929.8494 | 11.0386 | 100 | BMB* | 56.07%
31 | G2_Sample 5 | 1.805 | 43.3447 | 935.6866 | 11.1685 | 100 | BMB* | 57.97%
32 | G2_Sample 5 | 1.8033 | 43.4232 | 941.9338 | 11.1877 | 100 | BMB | 58.26%
33 | G2_Sample 5 | 1.8033 | 43.2816 | 944.3557 | 11.1531 | 100 | BMB | 57.74%
34 | G2_Sample 6 | 1.79 | 11.6535 | 244.7425 | 3.412 | 100 | BMB | 7.685%
35 | G2_Sample 6 | 1.7917 | 11.4713 | 239.8335 | 3.3674 | 100 | BMB* | 7.56%
36 | G2_Sample 6 | 1.7917 | 11.7905 | 243.6412 | 3.4455 | 100 | BMB | 7.779%
37 | G2_Sample 7 | 1.79 | 12.2285 | 243.4031 | 3.5527 | 100 | BMB* | 8.085%
38 | G2_Sample 7 | 1.7933 | 12.1989 | 237.4678 | 3.5455 | 100 | BMB | 8.064%
39 | G2_Sample 7 | 1.79 | 12.1822 | 238.1398 | 3.5414 | 100 | BMB | 8.052%
40 | G2_Sample 8 | 1.8033 | 43.3315 | 935.77 | 11.1653 | 100 | BMB* | 57.92%
41 | G2_Sample 8 | 1.8033 | 43.2818 | 928.5671 | 11.1531 | 100 | BMB* | 57.74%
42 | G2_Sample 8 | 1.8033 | 42.154 | 892.9663 | 10.8771 | 100 | BMB* | 53.82%
43 | G2_Sample 9 | 1.7917 | 12.1477 | 244.194 | 3.5329 | 100 | BMB* | 8.028%
44 | G2_Sample 9 | 1.79 | 11.6739 | 238.307 | 3.4169 | 100 | BMB* | 7.699%
45 | G2_Sample 9 | 1.79 | 12.0387 | 236.839 | 3.5063 | 100 | BMB* | 7.952%
46 | G2_Sample 10 | 1.8033 | 43.3194 | 918.4197 | 11.1623 | 100 | BMB* | 57.88%
47 | G2_Sample 10 | 1.805 | 43.2859 | 924.0726 | 11.1541 | 100 | BMB* | 57.76%
48 | G2_Sample 10 | 1.8033 | 43.5871 | 915.3108 | 11.2279 | 100 | BMB* | 58.87%
49 | G2_Sample 11 | 1.79 | 11.9117 | 233.1165 | 3.4752 | 100 | BMB* | 7.863%
50 | G2_Sample 11 | 1.7917 | 11.5806 | 231.0134 | 3.3941 | 100 | BMB* | 7.635%
51 | G2_Sample 11 | 1.7917 | 11.9222 | 231.1718 | 3.4777 | 100 | BMB* | 7.871%
52 | G2_Sample 12 | 1.8033 | 42.5759 | 913.0815 | 10.9803 | 100 | BMB* | 55.24%
53 | G2_Sample 12 | 1.8033 | 42.8231 | 909.4521 | 11.0409 | 100 | BMB* | 56.1%
54 | G2_Sample 12 | 1.805 | 43.9811 | 915.9308 | 11.3243 | 100 | BMB* | 60.36%
55 | G3_Sample 1 | 1.7883 | 11.5277 | 237.6883 | 3.3812 | 100 | BMB | 7.598%
56 | G3_Sample 1 | 1.79 | 11.5254 | 236.9967 | 3.3806 | 100 | BMB* | 7.597%
57 | G3_Sample 1 | 1.7883 | 11.0238 | 232.3682 | 3.2579 | 100 | BMB* | 7.256%
58 | G3_Sample 2 | 1.79 | 11.1301 | 221.9162 | 3.2839 | 100 | BMB* | 7.328%
59 | G3_Sample 2 | 1.79 | 11.0946 | 219.0985 | 3.2752 | 100 | BMB* | 7.304%
60 | G3_Sample 2 | 1.79 | 11.0749 | 217.4114 | 3.2703 | 100 | BMB* | 7.29%
61 | G3_Sample 3 | 1.8033 | 40.089 | 856.8893 | 10.3717 | 100 | BMB* | 47.47%
62 | G3_Sample 3 | 1.8017 | 41.0351 | 886.6319 | 10.6032 | 100 | BMB* | 50.26%
63 | G3_Sample 3 | 1.8033 | 41.2922 | 888.0547 | 10.6662 | 100 | BMB* | 51.05%
64 | G3_Sample 4 | 1.7867 | 10.4641 | 214.2783 | 3.1208 | 100 | BMB | 6.883%
65 | G3_Sample 4 | 1.7883 | 10.5373 | 211.3201 | 3.1388 | 100 | BMB | 6.931%
66 | G3_Sample 4 | 1.7883 | 10.4822 | 208.6858 | 3.1253 | 100 | BMB | 6.895%
67 | G3_Sample 5 | 1.8017 | 41.8807 | 872.8294 | 10.8102 | 100 | BMB* | 52.92%
68 | G3_Sample 5 | 1.8017 | 42.0995 | 884.7996 | 10.8637 | 100 | BMB* | 53.64%
69 | G3_Sample 5 | 1.8017 | 42.322 | 879.4179 | 10.9182 | 100 | BMB* | 54.38%
70 | G3_Sample 6 | 1.7883 | 10.8374 | 214.7623 | 3.2122 | 100 | BMB | 7.131%
71 | G3_Sample 6 | 1.7883 | 10.8871 | 217.3839 | 3.2244 | 100 | BMB | 7.164%
72 | G3_Sample 6 | 1.7867 | 10.7921 | 215.8892 | 3.2011 | 100 | BMB* | 7.101%
73 | G3_Sample 7 | 1.8017 | 42.2605 | 884.6074 | 10.9032 | 100 | BMB* | 54.17%
74 | G3_Sample 7 | 1.8033 | 42.4707 | 886.0778 | 10.9546 | 100 | BMB* | 54.88%
75 | G3_Sample 7 | 1.8 | 42.5322 | 900.9024 | 10.9697 | 100 | BMB* | 55.09%
76 | G3_Sample 8 | 1.7883 | 11.5588 | 231.1134 | 3.3888 | 100 | BMB | 7.62%
77 | G3_Sample 8 | 1.7883 | 11.4128 | 228.4007 | 3.3531 | 100 | BMB | 7.52%
78 | G3_Sample 8 | 1.7867 | 11.2988 | 227.6187 | 3.3252 | 100 | BMB* | 7.442%
79 | G3_Sample 9 | 1.8 | 42.4392 | 883.3276 | 10.9469 | 100 | BMB* | 54.78%
80 | G3_Sample 9 | 1.8017 | 42.0231 | 886.6521 | 10.8451 | 100 | BMB | 53.39%
81 | G3_Sample 9 | 1.8 | 41.9322 | 879.1608 | 10.8228 | 100 | BMB* | 53.09%
82 | G3_Sample 10 | 1.7867 | 11.775 | 240.3877 | 3.4417 | 100 | BMB | 7.768%
83 | G3_Sample 10 | 1.7883 | 11.9626 | 242.9828 | 3.4876 | 100 | BMB* | 7.899%
84 | G3_Sample 10 | 1.7883 | 11.9317 | 239.5623 | 3.4801 | 100 | BMB | 7.877%
85 | G3_Sample 11 | 1.8 | 41.8542 | 878.5256 | 10.8037 | 100 | BMB | 52.83%
86 | G3_Sample 11 | 1.8 | 42.8064 | 897.0727 | 11.0368 | 100 | BMB | 56.04%
87 | G3_Sample 11 | 1.8017 | 42.9452 | 895.4426 | 11.0707 | 100 | BMB* | 56.53%
88 | G3_Sample 12 | 1.7867 | 12.0807 | 241.8581 | 3.5165 | 100 | BMB | 7.981%
89 | G3_Sample 12 | 1.7867 | 11.9881 | 238.789 | 3.4939 | 100 | BMB | 7.916%
90 | G3_Sample 12 | 1.79 | 11.9523 | 236.6184 | 3.4851 | 100 | BMB | 7.891%

## Calibration
Calibration
Calibration Details |  | IPA
Calibration Type |  | Lin, WithOffset, Avg |  |  | Offset (C0) |  | -2.2869
Evaluation Type |  | Area |  |  | Slope (C1) |  | 4.0857
Number of Calibration Points |  | 10 |  |  | Curve (C2) |  | 0
Number of disabled Calibration Points |  | 0 |  |  | R-Square |  | 0.99948
Calibration Plot |  | IPA
Calibration Results |  | IPA
No. | Injection Name | Calibration | X Value | Y Value | Y Value | Area | Height
|  | Level |  |  |  | pA*min | pA
|  |  | FID | FID | FID | FID | FID
|  |  | IPA | IPA | IPA | IPA | IPA
4 | Standard solution IPA S1 - 2.58M | 1 | 2.58 | 8.6425 | 8.6425 | 8.6425 | 206.4439
5 | Standard solution IPA S2 -3.93M | 2 | 3.93 | 13.7896 | 13.7896 | 13.7896 | 327.2576
6 | Standard solution IPA S3-5.16M | 3 | 5.16 | 18.8365 | 18.8365 | 18.8365 | 447.4335
7 | Standard solution IPA S4-6.53M | 4 | 6.53 | 24.1656 | 24.1656 | 24.1656 | 567.6417
8 | Standard solution IPA S5-7.71M | 5 | 7.71 | 29.1615 | 29.1615 | 29.1615 | 676.3403
9 | Standard solution IPA S6-9.11M | 6 | 9.11 | 34.4926 | 34.4926 | 34.4926 | 793.8218
10 | Standard solution IPA S7-10.36M | 7 | 10.36 | 39.5562 | 39.5562 | 39.5562 | 912.1902
11 | Standard solution IPA S8-11.67M | 8 | 11.67 | 45.296 | 45.296 | 45.296 | 1020.839
12 | Standard solution IPA S9-12.44M | 9 | 12.44 | 49.0797 | 49.0797 | 49.0797 | 1114.0516
13 | Standard solution IPA S10-12.9M | 10 | 12.9 | 50.7344 | 50.7344 | 50.7344 | 1132.6117

## Chromatogram
Chromatogram

## Audit Trail
Audit Trail
Injection Details
Injection Name: |  | Standard solution IPA S10-12.9M | Run Time (min): | 3.6983
Vial Number: |  | 2 | Injection Volume: | 0.5
Injection Type: |  | Calibration Standard | Channel: | FID
Calibration Level: |  | 10 | Wavelength: | n.a.
Instrument Method: |  | IPA Stabilwax 2024-02-20_high conc | Bandwidth: | n.a.
Processing Method: |  | IPA_2024-02-19_high conc | Dilution Factor: | 1
Injection Date/Time: |  | 2025-02-27 11:31 | Sample Weight: | 1
Day Time | Ret. Time | Command/Message
(UTC+00:00) | min
11:26:03 |  | Audit trail of injection "chrom://md-051298/ChromeleonLocal/Teaching methods/2025-EX3502 Lab/2025-02-27 10 am EX3502.seq/976.smp". Executed 27/02/2025 on instrument TRACE1300 (instrument controller MD-051298).
11:26:03 |  | Start of injection 13 "Standard solution IPA S10-12.9M", using instrument method "IPA Stabilwax 2024-02-20_high conc".
11:26:03 | 0.000 | Entered stage "Instrument Setup"
11:26:03 | 0.000 | GC.SSL.FlowMode = FlowCtrl
11:26:03 | 0.000 | Sampler.DrawSpeed = Slow
11:26:03 | 0.000 | Sampler.FillStrokes = 3
11:26:03 | 0.000 | Sampler.AirVolume = 1.00 µl
11:26:03 | 0.000 | Sampler.GcType = TRACE_1300_1310
11:26:03 | 0.000 | Sampler.PostWash = 6
11:26:03 | 0.000 | Sampler.PostWashVial = D
11:26:03 | 0.000 | Sampler.SampleWash = 3
11:26:03 | 0.000 | Sampler.PreWash = 2
11:26:03 | 0.000 | Sampler.PreWashVial = B
11:26:03 | 0.000 | GC.PrepRunTimeout = 999.99 min
11:26:03 | 0.000 | GC.EquilibrationTime = 0.50 min
11:26:03 | 0.000 | GC.ReadyDelay = 0.00 min
11:26:03 | 0.000 | GC.SSL.SplitMode = Split
11:26:03 | 0.000 | GC.SSL.SplitFlow.Nominal = 240.0 ml/min
11:26:03 | 0.000 | GC.SSL.SplitFlowCtrl = On
11:26:03 | 0.000 | GC.SSL.PurgeFlow.Nominal = 5.000 ml/min
11:26:03 | 0.000 | GC.SSL.PurgeFlowCtrl = On
11:26:03 | 0.000 | GC.SSL.VacuumCorrection = Off
11:26:03 | 0.000 | GC.SSL.GasSaverFlow = 10.0 ml/min
11:26:03 | 0.000 | GC.SSL.GasSaverTime = 2.00 min
11:26:03 | 0.000 | GC.SSL.GasSaverCtrl = On
11:26:03 | 0.000 | GC.SSL.Temperature.Nominal = 180 °C
11:26:03 | 0.000 | GC.SSL.TempCtrl = On
11:26:03 | 0.000 | GC.FrontColumn.Description = "Stabilwax-DA SN1645029"
11:26:03 | 0.000 | GC.FrontColumn.Length = 30.00 m
11:26:03 | 0.000 | GC.FrontColumn.NominalID = 0.250 mm
11:26:03 | 0.000 | GC.FrontColumn.FilmThickness = 0.25 µm
11:26:03 | 0.000 | GC.FID.Data_Collection_Rate = 10 Hz
11:26:03 | 0.000 | GC.FID.AirFlow.Nominal = 350.0 ml/min
11:26:03 | 0.000 | GC.FID.AirFlowCtrl = On
11:26:03 | 0.000 | GC.FID.MakeupFlowCtrl = Off
11:26:03 | 0.000 | GC.FID.H2Flow.Nominal = 35.0 ml/min
11:26:03 | 0.000 | GC.FID.H2FlowCtrl = On
11:26:03 | 0.000 | GC.FID.Temperature.Nominal = 260 °C
11:26:03 | 0.000 | GC.FID.TempCtrl = On
11:26:03 | 0.000 | GC.FID.Flame = On
11:26:03 | 0.000 | GC.FID.FlameoutRetry = On
11:26:03 | 0.000 | GC.FID.LitOffset = 1.0 pA
11:26:03 | 0.000 | GC.FID.PeakWidth = Standard
11:26:03 | 0.000 | GC.SSL.FlowCtrl = On
11:26:03 | 0.000 | GC.TempCtrl = On
11:26:03 | 0.000 | Entered stage "Equilibration"
11:26:03 | 0.000 | GC.SSL.Flow.Nominal = 1.200 ml/min
11:26:03 | 0.000 | GC.Temperature.Nominal = 40.0 °C
11:26:03 | 0.000 | Entered stage "Inject Preparation"
11:26:03 | 0.000 | GC.FID.Autozero
11:26:03 | 0.000 | Wait Sampler.Ready And GC.Ready
11:28:04 | 0.000 | New detailed run state: Waiting for equilibration time.
11:28:34 | 0.000 | Log RunState: StandBy
11:28:34 | 0.000 | Prep-run command sent to GC.
11:28:34 | 0.000 | Log RunState: PrepRun
11:28:34 | 0.000 | New detailed run state: Executing prep-run events
11:28:55 | 0.000 | Log RunState: Ready
11:28:55 | 0.000 | Log Ready: Ready
11:28:55 | 0.000 | Wait finished
11:28:55 | 0.000 | Entered stage "Inject"
11:28:55 | 0.000 | Sampler.Inject
11:28:55 | 0.000 | Injecting from vial position 2.
11:28:55 | 0.000 | Injection Volume is 0.5 µl.
11:28:55 | 0.000 | Waiting for inject response on Sampler.
11:28:55 | 0.000 | New detailed run state: Waiting for start key
11:31:01 | 0.000 | Got inject response.
11:31:01 | 0.000 | Entered stage "Start Run"
11:31:01 | 0.000 | GC.FID.AcqOn
11:31:01 | 0.000 | Entered stage "Run"
11:31:01 | 0.004 | New detailed run state: Running
11:31:01 | 0.004 | Log RunState: Run
11:31:01 | 0.004 | Log Ready: NotReady
11:32:01 | 1.000 | GC.Temperature.Nominal Gradient Start = 40.0 °C, End = 125.0 °C, Duration = 1.700 min
11:33:43 | 2.700 | GC.Temperature.Nominal = 125.0 °C
11:34:43 | 3.700 | GC.Temperature.Nominal = 125.0 °C
11:34:43 | 3.700 | Entered stage "Stop Run"
11:34:43 | 3.700 | GC.FID.AcqOff
11:34:43 | 3.700 | GC.SSL.Flow.Nominal = 1.200 ml/min
11:34:43 | 3.700 | Terminating instrument method chrom://md-051298/ChromeleonLocal/Teaching methods/2025-EX3502 Lab/2025-02-27 10 am EX3502.seq/IPA Stabilwax 2024-02-20_high conc.instmeth. Waiting for instrument to finish processing of current injection ...
11:34:48 | 3.790 | New detailed run state: Cooling down
11:34:48 | 3.790 | Log RunState: Cool
11:34:48 | 3.795 | Log Temperature.Nominal: 40.0 °C
11:34:54 | 3.890 | Log RunState: NotReady
11:34:54 | 3.890 | New detailed run state: Waiting for ready
11:34:54 | 3.891 | OK to terminate instrument method chrom://md-051298/ChromeleonLocal/Teaching methods/2025-EX3502 Lab/2025-02-27 10 am EX3502.seq/IPA Stabilwax 2024-02-20_high conc.instmeth.
11:34:54 |  | End of injection "Standard solution IPA S10-12.9M".

## Standard solutions
2025-02-24
| Total Volume = 5 ml |  | Exact masses
Standard solution name | V(IPA), ml | V(H2O), ml | m (IPA), g | m (H2O), g | n (IPA), mol | n (H2O), mol | % Mol | C(IPA), Mol/L
Standard solution IPA S1 | 1 | 4 | 0.7757 | 4.0047 | 0.012907 | 0.22248 | 5.483% | 2.5814
Standard solution IPA S2 | 1.5 | 3.5 | 1.1814 | 3.5057 | 0.019657 | 0.19476 | 9.168% | 3.9314
Standard solution IPA S3 | 2 | 3 | 1.5518 | 3.0111 | 0.02582 | 0.16728 | 13.37% | 5.1641
Standard solution IPA S4 | 2.5 | 2.5 | 1.9612 | 2.5184 | 0.032632 | 0.13991 | 18.91% | 6.5265
Standard solution IPA S5 | 3 | 2 | 2.3177 | 2.0137 | 0.038564 | 0.11187 | 25.63% | 7.7128
Standard solution IPA S6 | 3.5 | 1.5 | 2.7382 | 1.5146 | 0.045561 | 0.084144 | 35.13% | 9.1121
Standard solution IPA S7 | 4 | 1 | 3.1122 | 0.9931 | 0.051784 | 0.055172 | 48.42% | 10.3567
Standard solution IPA S8 | 4.5 | 0.5 | 3.5072 | 0.4961 | 0.058356 | 0.027561 | 67.92% | 11.6712
Standard solution IPA S9 | 1.8 | 0.2 | 3.7392 | 0.1975 | 0.062216 | 0.010972 | 85.01% | 12.4433
Standard solution IPA S10 | 5 | 0 | 3.8769 | 0 | 0.064507 | 0 | 100% | 12.9015
MW IPA, g/mol | MW H2O, g/mol | IPA Density @ 20°C, g/mL | Water Density @ 20°C, g/mL
60.1 | 18 | 0.786 | 0.9982